## Slide 1
Author E-mail: drfcabib@mps.com.br
SAMPLING ERROR OF KONAN SPECULAR MICROSCOPE CELL CHECK XL AND NIDEK CEM 530
Fernando Cesar Abib, MD, PhD *
Federal University of Parana
Hospital Erasto Gaertner, Brazil
Dulcemar Szeremeta Abib, MD
Federal University of Parana, Brazil
Eposter 90
Notes: 1

## Slide 2: Purpose
Get to know the existing sampling error in corneal Specular Microscopy (SM) examinations performed routinely with the appliances Konan Cell Check XL and Nidek CEM530.

## Slide 3
Methods
Retrospective study
  30 corneal SM performed with Konan Cell Check XL;
  30 corneal SM performed with Nidek;
The number of endothelial cells counted in each examination, along with the results obtained in specular microscopy will be analyzed by the CELLS ANALYZER software - Corneal Endothelial Statistical Lab to the reliability degree of confidence of 95% and the relative error of 0.05 (5%);
The relative error is called the sampling error and the examination of the sample will be considered sufficient when the sampling error calculated is less than 5%, that is, accepted by the medical science.

## Slide 4: Results
Konan Cell Check XL
  Age of patients ranged from 14 to 77 years old with an average of 47 ± 16 years;
  Characterization of the examined endothelium:
    Endothelial average density 2686 ± 372 cells/mm²;
    Coefficient of variation 0.370 ± 0.120;
    Percentage of hexagonal cells 44.2 ± 8.1%;
  Characterization of endothelial sample:
    Number of cells counted in surveys ranged 91-181 averaging 139 ± 19 cells;
    Sampling error calculated ranged from 0.040 to 0.1534 averaging 0.060 ± 0.021;
    Only 7 of the 30 examinations (23%) had in their performance, sufficient cell numbers to represent the reality of the endothelial mosaic to the confidence level of 95% and a sampling error of less than 5% error accepted by the medical science.
Nidek CEM 530
  Age of patients ranged from 23 to 71 years old with an average of 52 ± 12 years;
  Characterization of the examined endothelium:
    Endothelial average density 2361 ± 286 cells / mm²;
    Coefficient of variation 0.335 ± 0.053;
    Percentage of hexagonal cells 56.2 ± 9%;
  Characterization of endothelial sample:
    Number of cells counted in surveys ranged 63-222 averaging 140 ± 36 cells;
    Calculated sampling error ranged from 0.0335 to 0.1038 averaging 0.057 ± 0.015; Only 9 of the 30 examinations (30%) had in their performance, sufficient cell numbers to represent the reality of the endothelial mosaic to the confidence level of 95% and less than 5% sampling error, error accepted by the medical science.

## Slide 5: Conclusion
Most of the performed corneal Specular Microscopy did not include sufficient number of cells to represent the reality of the examined corneal endothelium;
It is suggested that the Reliability Indexes used in this study are routinely adopted to calculate the sample size (number of cells counted) required for each Specular Microscopy.